MAST AND DISDROMETER DATA: FILES, DATA STRUCTURE AND INSTRUMENTATION

Description: Data in this directory are measured on the Risø site (55 41 39.1 N,12 05 17.7 E in WGS 84), where a 122m met mast is located 6.5m above local sea level. The site is located on the coast of the Roskilde fjord, on the Danish island of Sjælland. The disdrometer was placed near the bottom of the tower.

Creator: Ebba Dellwik, ebde@dtu.dk

License: CC BY 4.0

Description: There is one file for the 10 minute mean values from the Risø mast observations, another file for the corresponding standard deviations and a third containing disdrometer data. All files include data in ascii format.

MAST DATA: QUALITY CONTROL AND DATA STRUCTURE

The mast data has been quality controlled by visual inspection and obvious problems have been marked with -9999. These problems include electrical noise which affect the standard deviation of the wind direction measured at 77m after Dec 16th 2023, and faulty data of RH, radiation, temperature and pressure data before February 3rd 2023.

Cup anemometers and vanes were replaced in March 2023 since the calibration certificates were outdated and there may, therefore, be minor changes to the instruments' sensitivity and offsets prior to this date.

The instruments are mounted on booms which point towards the direction of 225 deg. The instruments are, therefore, in the wind shadow of the mast in the direction of 45 deg. For wind directions of 45 pm 20 degrees, cup anemometer readings are artificially low compared to the undisturbed flow, and, correspondingly, artificially high for the standard deviations. Depending on the need of accuracy, data from this wind direction interval should be discarded.

| File_name | Description | Size | # Files | Dates |
| --- | --- | --- | --- | --- |
| Risoe_mast_data_AIRE_one_year_mean.asc | Mean values of mast observations | 16.1MB | 1 | 2023-01-01- - 2024-02-15 |
| Risoe_mast_data_AIRE_one_year_std.asc | Standard deviations of mast observations | 16.1MB | 1 | 2023-01-01- - 2024-02-15 |
| Parsivel2_Risoe_Site_10minutes_RAW.csv | Disdrometer data, not quality controlled | 91.5MB | 1 | 2023-03-17 – 2024-02-15 |

The two files with mast data contain the following measurements. The header contains the field_name, as listed below.

| Field_name | Description | Variable | Unit |
| --- | --- | --- | --- |
| Year month day hour minute | Date-Time in UTC+1, end of 10 minute measurement period. | Time |  |
| Cup44m | Risø cup P2546D at 44m above local ground | Wind speed | m/s |
| Cup77m | Risø cup P2546D at 77m above local ground | Wind speed | m/s |
| Cup94m | Risø cup P2546D at 94m above local ground | Wind speed | m/s |
| Cup118m | Risø cup P2546D at 118m above local ground | Wind speed | m/s |
| Cup125m | Risø cup P2546D at 125m above local ground | Wind speed | m/s |
| Vane77m | F2919A W200P vane, Vector | Wind direction | º |
| Vane94m | F2919A W200P vane, Vector | Wind direction | º |
| RainFromRainGauge | RIM-7499-020-STD, RIMCO | Rain amount | mm/10min |
| Pressure2m | Vaisala PTB 110 | Air pressure | hPa |
| RH2m | Vaisala HMP155 | Relative humidity | % |
| T2p5m (deg C) | Risø P2449A temperature sensor with Risø P2029 radiation shield | Air temperature | º C |
| Radiation_SW_120m | CM11 Kipp&Zonen Pyranometer | Short wave incoming radiation | W/m2 |

DISDROMETER DATA

One year of 10 minute mean values from the the Parsivel2 sensor from OTT can also be found here. Please consult the manual for information on data structure. A quality-controlled dataset will be uploaded within the next few months.